## Slide 1: Evaluation of Machine Learning Algorithms for Multi-parameter Patient Monitoring
Mohammad Adibuzzaman
ORISE Fellow, FDA
PhD Candidate, Marquette University

## Slide 2: Outline
Introduction
Study design and dataset
Approaches (decision tree and support vector machine, early warning score)
Results (sensitivity and specificity)
Summary

## Slide 3: Introduction
Poor interpretation of patient state
Alarm fatigue
Heart Rate
Blood Pressure
Oxygen Saturation
Data Fusion Algorithms
Automatic Warnings
Timely Clinical Response (e.g. medical emergency team (MET) activation)

## Slide 4: Introduction
Early warning systems include algorithms that use multiple vital signs to monitor patients and recognize early deterioration.
Performance evaluation of these algorithms are critical to provide warnings with high sensitivity and reduced numbers of false alarms to address alarm fatigue.
The aim of this research is to investigate how algorithm development techniques (selection of training, testing and validation dataset) can affect performance using a publicly available dataset.

## Slide 5: Regulatory Impact
Multiple submissions have come in using these approaches
  Visensia (FDA cleared data fusion device for monitoring of patient deterioration)
How to evaluate the machine learning algorithms?
What data set should be used for training, testing and validation?
Are the machine learning algorithms ‘better’ than existing reference early warning scoring systems?

## Slide 6: Hypothesis
We hypothesize that the performance of machine learning algorithms to predict critical hospital events improves when the algorithm is trained on larger data sets.

## Slide 7: Data: Multi-parameter Intelligent Monitoring in Intensive Care (MIMIC II)
Clinical Database
Waveform Database
Matched Subset
~6,000 records
~28,000 records
~24,000 records
Intellivue MP-70 (Philips Healthcare)
Waveforms
    ECG
    Blood pressure
    Plethysmography
Vital Signs
    Heart rate
    Blood pressure
    Oxygen saturation
    Cardiac output
Nurse entered physiological data
Medications
Laboratory data
Nursing notes
Discharge notes
CareVue Clinical Information System (Philips Healthcare)
Saeed et al. Multiparameter Intelligent Monitoring in Intensive Care II: a public-access intensive care unit database. Crit Care Med. 39:5. 2011.

## Slide 8: Data for Training, Testing and Validation
2400 records from the matched subset of Multi-parameter Intelligent Monitoring in Intensive Care (MIMIC2) with 5 hours of continuous trend data.
2400 records with 5 hours of trend data
500-1
500-2
400
500-3
500-4
1000-1
400
1000-2
1500-1
400
4 sets of 500 records
2 sets of 1000 records
1 set of 1500 records
Validation
2000-1
400
Train and Test
1 set of 2000 records
Train
Test
Train
Test
Test
Train
Test
10 fold cross validation for each set
500
500
500
Select the best performance algorithm, and test that with the independent validation set

## Slide 9: Designing an Algorithm to Predict Clinical Events
Observation Window
(60 Min)
Gap Window
(60 Min)
Target Window
(10 Min)
Critical hospital events e.g. Medical Emergency Team (MET) Activation Criteria

## Slide 10: Medical emergency team activation (MET) criteria
Medical emergency team (MET) activation criteria is used in hospital settings as ‘clinically critical events’.
MET activation criteria was set as the event to predict.
MET activation criteria
  Heart rate (<40 or >160 bpm)
  Systolic blood pressure(<60 or >200 mmHg) or
  Oxygen saturation (<85%)
The goal is to predict if MET activation will occur 60 minutes in advance (in the target window).

## Slide 11: Algorithm
Two machine learning approaches (decision tree and support vector machine SVM) were applied to develop the MET prediction algorithm using three vital signs (heart rate, systolic blood pressure and oxygen saturation).
Features (Mean, median, variance, kurtosis, skewness, cross correlation and early warning score)
Heart Rate
Oxygen Saturation
Observation window containing vital signs
Systolic Blood Pressure
Predict MET activation 60 minutes in advance
MET activation prediction in the target window
Machine Learning Algorithms

## Slide 12: True Positive Event: Medical Emergency Team Activation (MET) Criteria Satisfied
Oxygen Saturation drops below 85%
(MET activation criteria)
Observation
window
Gap
window
Target window

## Slide 13: True Negative Event: Medical Emergency Team Activation (MET) Criteria not Satisfied
Observation
window
Gap
window
Target window
Target window does not satisfy MET activation criteria

## Slide 14: Comparison Approaches
Decision tree (Machine learning algorithm)
Support vector machine (Machine learning algorithm)
National early warning score (NEWS) (Reference method)

## Slide 15: Decision Tree
Mean(Heart rate) >80
Yes
No
Mean(Systolic blood pressure) >200
Mean(Oxygen saturation) <85
Variance
(Oxygen saturation)
<0.4
Variance
(Systolic blood pressure)
> 0.8
MET
activation
No
MET
activation
MET
activation
No
MET
activation
MET
activation
No
MET
activation
Yes
No
Yes
No
Yes
No
Yes
No

## Slide 16: Support Vector Machine
_
Margin
Separating Hyperplane
+
+
+
+
+
+
+
+
_
_
_
_
_
_
+
+
_
Mean Heart Rate (Feature 1)
Mean Blood Pressure (Feature 2)

## Slide 17: Early Warning Scores: Reference System
PHYSIOLOGICAL PARAMETER | 3 | 2 | 1 | 0 | 1 | 2 | 3
Oxygen Saturation | ≤91 | 92-93 | 94-95 | ≥96 |  |  | 
Systolic BP | ≤90 | 91-100 | 101-110 | 111-219 |  |  | ≥220
Heart Rate | ≤40 |  | 41-50 | 51-90 | 91-110 | 111-130 | ≥131
National Early Warning Score (NEWS) in the UK
A positive event (prediction for MET activation) using NEWS
  If any of the vital signs reaches the threshold of 3 for 10 minutes.
  If the cumulative score reaches 5 for 10 minutes.

## Slide 18: Results: Sensitivity to Detect MET Activation an Hour Before It Occurs
National Early Warning Score (NEWS)
Each box plot represent the results for the 10 fold cross validation for the independent validation set.

## Slide 19: Results: Specificity to Detect MET Activation an Hour Before It Occurs

## Slide 20: Summary
With a fixed training set, there is much variability in the performance of the algorithms depending on the random split for training and testing.
Increasing the number of records in the training set did not necessarily increase algorithm sensitivity.
The ‘best’ algorithm using the test sets is not necessarily the best algorithm for an ‘independent validation set’.
Future work to improve early warning patient monitoring algorithms should investigate:
  - Key signal features to include in the algorithms.
  Alternative techniques to combine information from multiple sources.

## Slide 21: Extra Slide: Sensitivity Using the Test Set

## Slide 22: Extra Slide: Specificity Using the Test Set